Standard Operating Procedure (SOP)
Title: Lipid Extraction from Plasma or Serum Using the Modified Alshehry Method
Version: 1.1
Effective Date: July 18, 2025
Prepared by: Alan Maschek
Reviewed by: Alan Maschek
Applicable Instrumentation: LC-MS lipidomics

1. Purpose

This SOP outlines the procedure for extracting lipids from human or animal plasma using a single-phase butanol–methanol method as described by Alshehry et al. (2015). This protocol is optimized for LC-MS–based lipidomics applications and ensures compatibility with downstream quantitation and profiling workflows.

2. Safety Considerations

Always wear a lab coat, gloves, and eye protection.

Perform all steps involving volatile organic solvents in a certified chemical fume hood.

Dispose of waste solvents in accordance with institutional safety protocols.

3. Materials and Reagents

| Item | Details |
| --- | --- |
| Human or animal plasma | Preferably EDTA plasma; stored at –80°C |
| Methanol (LC-MS grade) | Store tightly sealed at RT |
| Butanol (LC-MS grade) | Store tightly sealed at RT |
| Milli-Q water | For dilution |
| Internal standard mixture (ISTD) | Custom or commercial lipid ISTD cocktail |
| Glass centrifuge tubes with Teflon caps | 15 mL |
| Microcentrifuge tubes | 1.5 or 2.0 mL |
| Vortex mixer |  |
| Benchtop centrifuge | Capable of 13,000 × g |
| Nitrogen evaporator or vacuum concentrator | With appropriate solvent trap |
| LC-MS vials | Amber recommended |

4. Sample Requirements

- Standard input: 10 µL plasma. Other volumes (5–50 µL) may be used with proportional solvent scaling. Record the total volume prior to extraction

5. Reagent Preparation

5.1 Extraction Solvent with Internal Standards

1:1 Butanol : Methanol (v/v)

• Mix equal volumes of butanol and methanol in a glass container.

• Store at RT in a tightly sealed glass bottle until internal standards are added.

Optional: Spike internal standards into the 1:1 butanol:methanol solvent or add separately during sample prep.

| Standard | Avanti PN | pmol/sample |
| --- | --- | --- |
| C16 Ceramide-d7 (d18:1-d7/16:0) | 860676P-1mg | 367.0 |
| C18 Ceramide-d7 (d18:1-d7/18:0) | 860677P-1mg | 349.0 |
| C24 Ceramide-d7 (d18:1-d7/24:0) | 860678P-1mg | 304.3 |
| C24:1 Ceramide-d7 (d18:1-d7/24:1) | 860679P-1mg | 305.3 |
| C18:1 Dihydroceramide | 860624P-5mg | 44.2 |
| C8 Dihydroceramide | 860626-5mg | 77.9 |
| N-12:0-1-deoxysphinganine | 860481-1mg | 35.6 |
| C17 Glucosylceramide | 860569P-5mg | 280.1 |
| Sphinganine-1-phosphate-d7 | 860694P-1mg | 34.3 |
| Sphingosine-1-phosphate-d7 | 860659P-1mg | 34.5 |
| Sphingosine-d7 | 860657P-1mg | 32.6 |
| 16:0-d31 SM | 868584 | 681.0 |
| PC(15:0-18:1(d7)) | 791637C-1mg | 663.9 |
| DG(15:0-18:1(d7)) | 791647C-1mg | 850.4 |
| TG(15:0-18:1(d7)-15:0) | 791648C-1mg | 615.5 |

- Store prepared solvent at −20 °C. Allow to equilibrate fully to room temperature and vortex before use.

- This internal standard panel is optimized for sphingolipid and glycerolipid quantitation and may be adjusted based on analytical scope. For broad untargeted lipidomics applications, a comprehensive mixture such as Avanti SPLASH® LIPIDOMIX® (Avanti Polar Lipids, Cat. No. 330707) is recommended. A typical starting amount is 5 µL of SPLASH per 20 µL plasma (0.25× sample volume), scaled proportionally to sample input. Consult lipidomics core for standard specifics.

6. Sample Preparation Procedure

6.1. Thawing

- Remove plasma samples from –80°C freezer and thaw on ice.

- Vortex briefly to ensure homogeneity.

6.2. Lipid Extraction

- Label tubes appropriately.

- Aliquot 10 µL of plasma into a 1.5 mL microcentrifuge tube.

- Add 100 µL of 1:1 butanol : methanol containing internal standards.

- Vortex for 10 - 30 seconds.

- Sonicate the samples in a water bath sonicator at 20 °C for 60 minutes. Ensure tubes are fully immersed and securely capped.

- Centrifuge at 13,000 × g for 10 minutes at 4 °C. This is a single-phase extraction; no water addition or phase separation step is required. Extracts may be injected directly or dried and reconstituted depending on analytical workflow.

- Carefully transfer the supernatant (~100 µL) while avoiding the precipitated protein pellet into a clean LC-MS vial or microcentrifuge tube.

6.3. Storage

- Store extracts at −80 °C for long-term storage. Short-term storage at −20 °C is acceptable; avoid repeated freeze–thaw cycles.

7. Notes

- The 1:10 (v/v) ratio of plasma to solvent is critical for reproducibility.

- A shorter incubation is acceptable, e.g. a 15 minute sonication has been shown to provide same coverage as 60 minute sonication.

- For larger volumes of plasma, scale proportionally and confirm extraction efficiency.

- Glassware should be solvent-rinsed and dried prior to use.

- Include blank and QC samples in every extraction batch.

8. Quality Control

- Include at least one pooled plasma QC per batch.

- Analyze a method blank (solvent only) alongside experimental samples.

- Monitor ISTD recovery and retention times throughout LC-MS analysis.

9. Troubleshooting

| Issue | Possible Cause | Solution |
| --- | --- | --- |
| Low recovery of lipids | Incomplete mixing or extraction | Verify vortex and incubation steps |
| Poor reproducibility | Solvent evaporation or pipetting errors | Use tight caps; calibrate pipettes |
| ISTD signal variation | Inconsistent spiking | Pre-mix ISTDs in extraction solvent |

10. Documentation

- Record sample ID, date, volume used, internal standard batch, and operator initials in the extraction log.

- Log deviations from SOP and corrective actions taken.

11. References

Alshehry, Z. H., et al. (2015). "An efficient single phase method for the extraction of plasma lipids." Metabolites, 5(3), 389–403. DOI: 10.3390/metabo5030389